## Slide 1: Trima Install Scripts
Installs new version of Trima SW on the machine
Checks versions to make sure it is compatible with the hardware
Converts config files between versions
  Checks ranges on config.dat items

## Slide 2: Source Files
Code is in ../Trima/Install/InstallCommon
  installer.h/cpp
  updatecassette.h/cpp
  updatecassette_dat.h/cpp
  updateTrima.h/cpp
  updateTrimaData.h
  updateTrimaDefines.h
  updateTrimaUtils.h/cpp
  /Install5X and /Install6X are the build areas for VxWorks 2.02 and 2.2 versions of the scripts

## Slide 3: updateTrima.h/cpp
Figures out what is installed, what is to be installed, what the hardware is, if install is allowed
Determines SW versions based on build numbers
Extracts source files
updateTrima() is the entry point from VxWorks
Calls installer::upgrade() with “from” and “to” version

## Slide 4: installer.h/cpp
Contains the code that installs new config files and modifies config.dat as needed
Runs conversion functions based on the new Trima version and data in updatetrimadata.h
For config.dat, it uses the new /templates/config.dat as a base and migrates user configured values to the new file

## Slide 5: updatetrimadata.h
List of the supported Trima builds and what conversion functions to run when installing them.
  Install will use the data for the closest build with an equal or smaller build number. (e.g. if build 6.327 is installed, the flags for 6.301 will be used)
  It won’t jump down major or minor rev numbers (e.g. if build 9.0.020 is installed and the smallest 9.0. build number is 9.0.034, it won’t use the flags for 8.9.00001)
  If Install can’t find an appropriate build, the build will fail

## Slide 6: updatetrimadata.h con’t
List of config.dat range values that are different across Trima versions for range checking.
  Settings that have the same ranges across all versions aren't included because the Trima software will enforce ranges prior to installing a new version.
Items in the PRODUCT_DEFINITIONS section will end in an underscore and all of the “a, b, c,..” products will be checked.

## Slide 7: Other files
updatetrimaUtils.h/cpp
  Utilities used by Install
updatetrimaDefines.h
  All of the #defines for Install
updatecassette.h/cpp & updatecassette_dat.h/cpp
  Cloned & slightly modified versions of these Trima utilities

## Slide 8: Adding New Trima Version
Unless there are major structural changes in the new Trima version, adding new lines to updatetrimadata.h is all that is needed.
If the new version is just a build number increment of an existing version and the checks & config.dat ranges are the same, nothing needs to be done. The data for the closest, smaller build will be used.

## Slide 9: Adding New Trima Version con’t
If conversion flags are different or there are config.dat range differences, a new entry in buildData will be needed.
If there are config.dat range differences, new entries in rangeData will be needed.
  The rangeType parameter in the buildData entry must correspond to a rangeType in the rangeData list.